## Slide 1
ΒΗΜΑΤΟΔΟΤΕΣ ΧΩΡΙΣ ΗΛΕΚΤΡΟΔΙΑ
(LEADLESS PACEMAKERS)
ΣΤΕΛΙΟΣ ΠΑΡΑΣΚΕΥΑÏΔΗΣ
ΔΙΕΥΘΥΝΤΗΣ ΕΣΥ
Α΄ Καρδιολογική Κλινική ΑΠΘ,
Νοσοκομείο ΑΧΕΠΑ, Θεσσαλονίκη

## Slide 2
NO CONFLICT OF INTEREST

## Slide 3: Pacemakers implants (2022)
pacemaker-related complications
Acute (2 months): 12 %
Chronic
lead or surgical pocket related: 10 %
> 1.500.000 /year worldwide, > 200.000 in USΑ
Udo E et al. Heart Rhythm 2012;9:728-735
The transvenous lead is a critical component of conventional cardiac pacemakers, but it is also their Achilles heel

## Slide 6
Lead fracture
Lead dislodgement

## Slide 7
1952, epicardial,
Paul Zoll

## Slide 9: History of Leadless Pacing: Idea proposed in 1970
1st self-contained leadless cardiac pacemaker (18 x 8 mm). Unreliable battery lasting weeks
Notes: The first totally self-contained leadless pacemaker system was proposed by Spickler back in 1970 using a device powered by mercury-zinc and nuclear power that was successfully tested in animals.
Source:
Spickler, William J., et al. Totally Self-Contained Intracardiac Pacemaker. J. Electrocardiology. 1970; 3 (3-4): 325-331.

## Slide 10: History of Leadless Pacing: Miniature Pacemaker Tested in 1991
a: Guiding catheter
b: Pusher catheter
c: Miniature pacemaker
d: Steering arm
5.8 x 23 mm
Notes: In 1991, Vardas et al created a cylindrical VOO pacing device specifically for an animal study. The device measured 5.8mm in diameter and 23 mm in length, and it consisted of three batteries and a CMOS timer. The authors stated that “if [a miniature pacing device] can be made programmable without an unreasonable increase in size, if the battery can be made sufficiently long lasting or externally rechargeable, if technology can rise to meet these challenges, then such a miniature device might one day take its place among regularly used pacemakers and, eventually, even supercede them.”
Source:
Vardas P et al, A miniature pacemaker introduced intravenously and implanted endocardially – preliminary findings from an experimental study. Eur.J.CPE. 1991. 1:27-30.

## Slide 11: LEADLESS PACEMAKER

## Slide 12: LEADLESS PACEMAKER
MICRA TPS-27 Fr sheath
NANOSTIM LCP- 18 Fr sheath
CE Mark 2015 / FDA APPROVAL 2016
FDA APPROVAL 2022

## Slide 13
Aveir DR i2i Study (ongoing)
RA
RV

## Slide 14
ESC GUIDELINES 2021

## Slide 15: αντενδειξεισ (fda)
μηχανική τριγλώχινα βαλβίδα
άλλες εμφυτευμένες καρδιακές συσκευές
νοσογόνος παχυσαρκία ( δυσχερής τηλεμετρία )
αλλεργία σε τιτάνιο, νιτινόλη, σκιαστικά-
ηπαρίνη / δυσανεξία στη δεξαμεθαζόνη
φίλτρα κάτω κοίλης φλέβας
ακατάλληλη ανατομία μηριαίας φλέβας
(διάμετρος, στένωση, απόφραξη, ελίκωση )
( Micra Transcatheter Pacing system – PRODUCT SPECIFICATION 2016 )

## Slide 16: SAME, STREAMLINED procedure
Implant tools
Micra Integrated Delivery Catheter
105 cm long catheter system with a handle that controls deflection and deployment of the Micra pacing capsule11
> 99%
Implant success3,4
Smooth Vessel Navigation with the Micra™ Introducer
Lubricious hydrophilic coating
23 Fr inner diameter (27 Fr outer diameter)
Silicone oil-coated dilator tip

## Slide 18: Testing fixation
“Pull and hold” test
2 tines have 15x the holding force necessary to hold the device in place.19

## Slide 21: LEADLESS PACEMAKER ΙΜPLΑΝΤΑΤΙΟΝ

## Slide 22: FLEXFIX NITINOL TINES
Multidimensional redundancy: Two tines have 15x the holding force necessary to hold the device in place19
Designed to minimize tissue trauma during deployment, repositioning, and retrieval21
Optimal electrode-tissue interface allows for low and stable chronic thresholds22
Low dislodgement rate (0.00-0.06%)3,4

## Slide 23: Testing electricals
Optimal electrode tissue interface allows for low and stable chronic thresholds.23

## Slide 24: Device lifecycle management options
Options
Micra™ transcatheter pacing systems can be turned OFF and additional Micra devices can be added.
      A Micra device takes up < 1% of the volume of a normal right ventricle.
Micra transcatheter pacing systems can be turned OFF and a traditional system or upgrade can be implanted.
The Micra design incorporates a proximal retrieval feature to enable acute retrieval.
Successful retrieval demonstrated after 4 years.
Proximal Retrieval Feature

## Slide 25: Multiple Micra™ Devices Can Be Implanted in the RV
Notes: Slide Notes:
The right ventricle can accommodate multiple Micra devices utilizing traditional pacing locations.

## Slide 26: Micra™ AV Transcatheter Pacing System VDD mode
The world’s smallest pacemaker, now with AV synchrony16
Micra AV’s accelerometer detects mechanical atrial activity and uses this information to deliver AV synchronous ventricular pacing.
New, integrated circuitry capable of sustaining new AV synchrony functionality.
Estimated average battery longevity of 8-13 years, dependent on the patient’s degree of AV block.
*AVB-only patients who would benefit from leadless pacing per the indications for use.
17
17

## Slide 27: Micra™ AV
ACCELEROMETER SIGNALS
A1 Start of ventricular systole, mitral and tricuspid valves close.
A2 End of ventricular systole, aortic and pulmonic valves close.
A3 Diastole, passive blood flow from A to V, corresponds to E-wave on Doppler echo.
A4 Atrial systole, blood pushed into ventricles, 100 ms electromechanical delay, corresponds to A-wave on Doppler echo.
Ventricular end (VE) marker
The end of the A1–A3 ventricular-event signals.
Atrial mechanical (AM) marker
Marker that indicates the device detected the atrial mechanical contraction or A4.

## Slide 28
MARVEL study- Micra Atrial TRacking Using A Ventricular
AccELerometer. VDD mode n=64
Heart Rhythm 2018
atrioventricular synchrony: 87%

## Slide 29: MICRA™ AV ALGORITHM PERFORMANCE JACC 2020
Results from the marvel 2 study
94.3%
Median AV synchrony at rest in complete AV block patients with normal sinus rhythm (n = 40).

## Slide 30: Primary results from the marvel 2 study18
MARVEL 2 in AV BLOCK PATIENT18
Note: MARVEL 2 clinical data slides were presented at AHA 2019.27

## Slide 31: MICRA TP study- Heart Rhythm 2018
n=1817 enrolled, 1801 (99%)successful implantations , 179 centers, 23 countries
follow up: 6.8 ± 6.9 months
3 cases of infection associated with the procedure, but none required device removal and there were no battery or telemetry issues

## Slide 32: Globally Diverse Patient Population
Two clinical studies spanning the globe41
2,543 patients, > 300 implanters, 208 centers, 31 countries
As of April 2018 data freeze

## Slide 33: Micra™ VR patients OLDER, MORE COMORBIDITIES41
 | Micra VR | Historical Control | 
 | (N = 726) | (N = 2,667) | p-value*
Age (years) | 75.9 ± 11.0 | 71.1 ± 12.1 | < 0.001
Female | 41.2% | 44.9% | 0.08
Hypertension | 78.7% | 67.2% | < 0.001
AF | 72.6% | 36.6% | < 0.001
Valvular Disease | 43.3% | 19.2% | < 0.001
Diabetes | 28.5% | 21.9%† | < 0.001
CAD | 28.2% | 38.4% | < 0.001
CHF | 18.0% | 15.0% | 0.050
COPD | 12.7% | 7.2%† | 0.001
Vascular Disease | 7.6% | 10.1% | 0.06
*P-value from T-test (continuous variables) or Fisher’s Exact test (categorical variables).
†Data parameter not collected across all 6 trials.

## Slide 34: Indications for pacing (N = 726)4
Chart: 
| Other | AVB | SND | Bradycardia with Permanent or Persistent AT/AF
 | 0.04 | 0.15 | 0.18 | 0.64
Consistent with Guideline Recommendations for VVI Pacing51
Reasons for selecting VVIR
Infrequent pacing expected
Advanced age
Sedentary lifestyle
Anatomical limitations
Comorbidities increasing complication risk
Other
4%
AVB 15%
SND 18%
Bradycardia with permanent or persistent AT/AF 64%

## Slide 35: Conditions That Precluded Use of a Transvenous System
Condition that Precluded Use of Transvenous System* | IDE Patients41 (N = 45) | PAR Patients41 (N = 435)
Compromised venous access | 34 | 134
Need to preserve veins for hemodialysis | 18 | 102
Thrombosis | 12 | 39
History of infection | 4 | 171
Cancer | 6 | 39
Other† | 6 | 138
*Not mutually exclusive.
†Including valve issues and prior issues with transvenous system.

## Slide 36: Procedure Times Reduced with Experience54
After 12 implants, median implant time reduced to 25.9 minutes.

## Slide 37: Implant site placement
Chart: 
| 1st Qtr | 2nd Qtr | 3rd Qtr | 4th Qtr
Sales | 0.66 | 0.33 | 0.01
Apex 66%
Septum 33%
RVOT 1%
Other 1%
NR 1%
IDE Site Placement (N = 720)41
Post-approval Site Placement (N = 1,801)41
Chart: 
| 1st Qtr | 2nd Qtr | 3rd Qtr | 4th Qtr
Sales | 32 | 64 | 1 | 1 | 1
Septum 64%
Apex 32%
RVOT 1%
Other
< 1%

## Slide 38: MAJOR COMPLICATIONS THROUGH 12 MONTHS Heart Rhythm 2018
Chart: Major Complication Rate through 12 Months (%)
| IDE (N = 726) | Post-market (N = 1,817)
 | 4 | 2.7

## Slide 39: major complication rate: 2.7 %
The risk of major complications for Micra PAR (Post Approval Registry) pts was 63% lower compared with transvenous pacemakers through 12 months F/U
Micra
Transvenous

## Slide 40: Micra™ VR major complications (N = 726)13
 | Within 30 days | 30 days– 6 months | > 6 months | Events (patients, %)
Total | 24 | 6 | 2 | 32 (29, 4.0%)
Cardiac Perforation/Effusion | 10 | 1 | 0 | 11 (11, 1.5%)
AV Fistula/Pseudoaneurysm | 5 | 0 | 0 | 5 (5, 0.7%)
Cardiac Failure | 0 | 4 | 2 | 6 (6, 0.8%)
Elevated Thresholds | 2 | 0 | 0 | 2 (2, 0.3%)
Pacemaker Syndrome | 1 | 1 | 0 | 2 (2, 0.3%)
Acute MI | 1 | 0 | 0 | 1 (1, 0.1%)
Deep Vein Thrombosis | 1 | 0 | 0 | 1 (1, 0.1%)
Metabolic Acidosis | 1 | 0 | 0 | 1 (1, 0.1%)
Presyncope | 1 | 0 | 0 | 1 (1, 0.1%)
Pulmonary Embolism | 1 | 0 | 0 | 1 (1, 0.1%)
Syncope | 1 | 0 | 0 | 1 (1, 0.1%)

## Slide 41: Low and stable pacing thresholds13
Battery Longevity Estimate Based on use conditions at 12 months, median battery longevity estimate is 12.1 years.*
Chart: PACING CAPTURE THRESHOLD
| Implant (n = 711) | Discharge (n = 640) | 1-month (n = 644) | 3-month (n = 685) | 6-month (n = 677) | 12-month (n = 630) | 18-month (n = 275) | 24-month (n = 58)
PCT | 0.62782 | 0.57834 | 0.57095 | 0.56926 | 0.59 | 0.6 | 0.58 | 0.53
*Use conditions included: median pacing 53.5%, median pacing threshold 0.50 V, median impedance 543 Ω; 89% of patients with > 10-year projected longevity; 99% of patients with > 5-year longevity.18
Mean ± standard deviation

## Slide 42: Micra™ VR electrical performance13
Chart: R-WAVE AMPLITUDE
| Implant (n = 690) | Discharge (n = 674) | 1-month (n = 653) | 3-month (n = 631) | 6-month (n = 609) | 12-month (n = 577) | 18-month (n = 254) | 12-month (n = 56)
Rwave | 11.1831 | 12.8208 | 14.8812 | 15.0602 | 15.1 | 15.1 | 15.4 | 15.5
Chart: PACING IMPEDANCE
| Implant (n = 720) | Discharge (n = 718) | 1-month (n = 698) | 3-month (n = 694) | 6-month (n = 682) | 12-month (n = 654) | 18-month (n = 284) | 24-month (n = 60)
Impedance | 723.597 | 678.579 | 643.754 | 619.683 | 606 | 596 | 614 | 619
Mean ± standard deviation
Mean ± standard deviation

## Slide 43
Observational
ESRD: end-stage renal disease
CCI: Charlson comorbidity index
EHJ 2022
n=16,431

## Slide 44
MICRA
MICRA
TRANSVENOUS
TRANSVENOUS
reinterventions
chronic complications
all-cause mortality

## Slide 45
EHJ 2022
+
-

## Slide 46
Lead-related Complications
Fractures
Insulation breaches
Venous thrombosis and obstruction
Tricuspid regurgitation

## Slide 47
JACC 2107
n=35, failed CRT, F/U: 6 months
n=12 with prior failed attempt for coronary sinus lead implantation
n=10 non responders to a previous CRT device
n=5, high CS pacing threshold or phrenic nerve capture

## Slide 48: WiSE-crt system – wireless LV endocardial pacing electrode

## Slide 49
HEART RHYTHΜ 2021

## Slide 50
MICRA + WiSE system
2021

## Slide 52: Πριν
ΜΕΤΑ

## Slide 53
2018
MICRA + S-ICD

## Slide 54: ΑΧΕΠΑ- ΒΗΜΑΤΟΔΟΤΕΣ ΧΩΡΙΣ ΗΛΕΚΤΡΟΔΙΑ – (LEADLESS PACEMAKERS)
Aυγουστος 2017- Δεκέμβριος 2022
20 εμφυτευσεις leadless pacemaker - 13 άνδρες, 7 γυναίκες
ηλικία: 60-87 (75 ± 6.3)

## Slide 55: HKΓ ΕΝΔΕΙΞΗ
Chart: 
| 1st Qtr | 2nd Qtr | 3rd Qtr
Column2 | 6 | 5 | 1

## Slide 56
Chart: 
| 1ο Τρ. | 2ο Τρ. | 3ο Τρ. | 4ο Τρ.
Πωλήσεις | 1 | 7 | 8 | 4
ΚΛΙΝΙΚΗ ΕΝΔΕΙΞΗ

## Slide 57: ΑΧΕΠΑ- ΒΗΜΑΤΟΔΟΤΕΣ ΧΩΡΙΣ ΗΛΕΚΤΡΟΔΙΑ
ΠΡΟΣΠΕΛΑΣΗ : δεξιά μηριαία φλέβα (100% )
ΘΕΣΗ ΕΜΦΥΤΕΥΣΗΣ ΣΥΣΚΕΥΗΣ : κορυφή δεξιάς κοιλίας n=12
μεσοκοιλιακό διάφραγμα, n=8
επανατοποθέτηση της συσκευής κατά την εμφύτευση : n=5, 25%)
- σε 3 ασθενείς απαιτήθηκε 1 επανατοποθέτηση
- σε 2 ασθενείς απαιτήθηκαν 2 επανατοποθετήσεις
ΜΕΣΗ ΔΙΑΡΚΕΙΑ ΕΠΕΜΒΑΣΗΣ : 80 ± 15 min → 50 ± 8 min
ΜΕΣΟΣ ΧΡΟΝΟΣ ΑΚΤΙΝΟΣΚΟΠΗΣΗΣ : 8,3 ± 1,8 min → 6 ± 1,2 min

## Slide 58: R WAVE
mean R wave ( implantation ) : 9 (±3,79) mV
mean R wave ( 30 days ) : 15,1 (±4,38) mV
ΟΥΔΟΣ ΒΗΜΑΤΟΔΟΤΗΣΗΣ
mean Capture Threshold ( implantation ) : 0,81 ± 0,48) V
mean Capture Threshold ( 30 days ) : 0,47 ±0,24 V
ΒΗΜΑΤΟΔΟΤΙΚΗ ΑΝΤΙΣΤΑΣΗ
mean Pacing Impedance ( implantation ) : 931 ±244 Ohm
mean Pacing Impedance ( 30 days ) : 680 ± 130 Ohm
επιπωματισμός : 1 κατά την εμφύτευση
ΕΠΙΠΛΟΚΕΣ

## Slide 59
παροδική κατακράτηση σκιαστικού ενδομυοκαρδιακά

## Slide 60
ΠΕΡΙΠΤΩΣΗ
διάβρωση θήκης βηματοδότη-εξωτερίκευση συσκευής

## Slide 61
Άνδρας 60 ετών
Συγγενής πλήρης ΚΚΑ
Εμφύτευση DDDR βηματοδότη- 1993 ( σε ηλικία 36 ετών )
Αντικατάσταση βηματοδότη ( 1996- 2003- 2010 )
Αρτηριακή υπέρταση (υπό αγωγή με μετοπρολόλη- φουροσεμίδη)
Υπερουριχαιμία
ΠΕΡΙΠΤΩΣΗ

## Slide 62: ECHO
Φυσιολογικές διαστάσεις καρδιακών κοιλοτήτων και ανιούσης αορτής
LVEF 50%, RV με διατηρημένη συστολική λειτουργία
Μικρού βαθμού MR, AR, TR
SPECT μυοκαρδίου- TC99m SESTAMIBI : φυσιολογικό
Αιματολογικες εξετάσεις
Αρνητικοί δείκτες φλεγμονής

## Slide 64
απύρετος
ΤΟΕ : χωρίς εκβλαστήσεις
καλλιέργειες αίματος και ούρων ( αρνητικες )
έναρξη εμπειρικής αντιβιοτικής αγωγής
βανκομυκίνη 1gr x 2 iv ( 4 εβδομάδες )
σιπροφλοξασίνη 400mg x 2 iv ( 4 εβδομάδες )
ριφαμπικίνη 300mg x 3 per os ( 2 εβδομάδες )

## Slide 65
αφαίρεση βηματοδότη και εξωαγγειακού τμήματος
ηλεκτροδίων
καθαρισμός θήκης και νεαροποίηση χειλέων θήκης
καλλιέργειες ( επίχρισμα, ιστοτεμάχια, καλώδια ):
αρνητικές
στόχος: εμφύτευση DDDR βηματοδότη δεξιά (γεννήτρια
και 2 νέα ηλεκτρόδια)

## Slide 67: ΑΠΟΦΡΑΞΗ ΑΝΩ ΚΟΙΛΗΣ ΦΛΕΒΑΣ
ΑΖΥΓΟΣ

## Slide 68
Rozmus, et al. J Interv Card Electrophysiol (2005) 13: 9
Venous thrombosis and stenosis (early or late ) 30- 50% of patients
1- 3% symptomatic ( including SVC syndrome )
Venous Thrombosis and Stenosis After Implantation of Pacemakers and Defibrillators

## Slide 69: θεραπευτικές επιλογές
χειρουργική αφαίρεση καλωδίων και εμφύτευση
επικαρδιακού βηματοδότη
βηματοδότης χωρίς ηλεκτρόδια - (leadless
pacemaker)

## Slide 72: ΤΟ ΜΕΛΛΟΝ ΘA EINAI ΑΣΥΡΜΑΤΟ ?

## Slide 73
ΕΥΧΑΡΙΣΤΩ ΓΙΑ ΤΗΝ ΠΡΟΣΟΧΗ ΣΑΣ